DUAL-CORE OPTICAL MOTION TESTBED MECHANICAL PLATFORM & INTERFACE PCB Rotary/translation hardware, differential encoder interface, STEP/DIR control, and physical verification Customer-facing technical report Revision 2.9 - 4 August 2026

# Executive Summary

This report presents the physical and electrical platform of the Dual-Core Optical Motion Testbed: the load-bearing rotary and translational platform, differential encoder receiver, STEP/DIR motor interface, and the firm-designed PCB with a dedicated solid-copper ground plane. The companion firmware report documents the Cortex-M7/M4 architecture, GUI, telemetry, safety behavior, and measured results.

| Prepared by | Herder Elektronische Systemen |
| --- | --- |
| Primary subject | Mechanical platform, custom PCB, encoder and motor-driver interfaces, and physical verification |
| Firmware context | STM32H747 Cortex-M7/M4 target with OpenAMP and shared-memory IPC |
| Distribution | Customer-facing; complete editable fabrication and private development archives excluded |

DOCUMENT PURPOSE This report is the physical and electrical half of the project. It leads with the custom interface PCB and mechanical platform, while the companion firmware FRD covers the dual-core target, GUI, telemetry, and results.

# 1. Platform and PCB Overview

The system combines a load-bearing rotary and translational optical platform with a firm-designed encoder/STEP-DIR interface PCB. The mechanical assembly provides controlled motion and optical mounting; the PCB provides differential encoder reception, logic-level feedback, motor commands, power/reference distribution, and a solid-copper ground plane for signal integrity.

| Figure 1. Exploded mechanical assembly and load path. | Figure 2. Populated interface PCB. |
| --- | --- |

| Subsystem | Purpose |
| --- | --- |
| Rotational stage | Continuous controlled rotation of the optical assembly |
| Translation stage | Manual lateral offset for alignment and crosstalk studies |
| Structural frame | Supports the optical breadboard and demonstrated static load |
| Interface PCB | Conditions encoder feedback and routes STEP/DIR control |
| Safety chain | Emergency-stop and target-side timeout behavior |

# 2. Mechanical Architecture

Figure 3. Exploded assembly showing the translation carriage, rotating platform, bearing structure, drive, and frame.

The lower structure supports the rotary axis and drive train. The upper carriage translates on parallel guides using a lead-screw mechanism, allowing the optical axis to be offset while preserving the load path through the frame.

| Mechanical requirement | Demonstrated or designed capability |
| --- | --- |
| Static load | 40 lb demonstrated |
| Translation | More than ±3 in |
| Operator-readable translation | Approximately 1/32 in |
| Rotating footprint | Approximately 12 × 12 in |
| Rotation | Continuous; operating speed depends on microstep and pulse configuration |

# 3. Assembled Platform and Translation

Figure 4. Assembled dual-axis optical motion platform.

Figure 5. High-contrast pointer and fixed scale for manual translation measurement.

# 4. Load and Safety Evidence

Figure 6. Static-load demonstration on the assembled platform.

The structure demonstrated the specified static load while retaining the rotary and translational architecture. This is prototype validation, not a formal structural or environmental certification.

Emergency-stop behavior removes drive power independently of the GUI.

Target-side deadman supervision stops motion after loss of host heartbeat.

Operator procedures must clear pinch points, moving cables, and optical hardware before motion.

The system requires a rigid optical table or equivalent mounting surface.

# 5. Encoder and STEP/DIR Interface Schematic

Figure 7. Featured encoder and STEP/DIR interface schematic.

The interface receives differential A/B encoder signals through the AM26LS32 line receiver, provides logic-level encoder feedback to the controller, and routes STEP/DIR control to the R725 motor driver. The controller, interface board, and driver share a defined 5 V logic rail and common ground reference.

| Connector | Signals / function |
| --- | --- |
| J1 | Differential encoder A+/A− and B+/B−, 5 V, GND |
| P1 | ENC_A, ENC_B, STEP_CTRL, DIR_CTRL, 5 V, GND to the controller |
| J2 | DIR, STEP, GND, 5 V to the motor-driver interface |

# 6. PCB Routing and Solid Ground Plane

| Figure 8. Routed signal layer and labeled interfaces. | Figure 9. Dedicated solid-copper ground plane. |
| --- | --- |

The redesigned board intentionally uses one layer as a continuous copper ground reference. Signal routing remains on the opposite layer wherever practical, reducing return-path discontinuities and making the reference strategy immediately reviewable.

| Design feature | Engineering purpose |
| --- | --- |
| Solid ground plane | Continuous low-impedance logic return path |
| Short receiver routing | Limits loop area between connector and AM26LS32 |
| Explicit connector labels | Reduces wiring and commissioning errors |
| Common logic reference | Maintains valid thresholds across controller and driver |
| Mounting clearances | Supports serviceable installation and connector access |

PUBLIC DESIGN EVIDENCE The customer-facing repository includes the schematic, routed-layer image, ground-plane image, and 3D renders. Editable fabrication files remain controlled engineering assets.

# 7. PCB Assembly and Component Placement

| Figure 10. Front PCB assembly view. | Figure 11. Angled PCB assembly view. |
| --- | --- |

| Designator / group | Function |
| --- | --- |
| U3 — AM26LS32 | Quad differential line receiver for encoder A/B |
| U1, U2 — NPN stages | STEP/DIR interface translation |
| J1 | Six-position encoder terminal block |
| J2 | Four-position motor-driver terminal block |
| P1 | Six-position controller header |

# 8. Controller, Encoder, and Drive Scaling

The PCB connects the mechanical plant to the STM32H747 target. Encoder feedback is captured through an interrupt-driven input path, and STEP/DIR commands are generated on the target. The physical scaling determines the relationship between firmware pulses, encoder counts, and table motion.

| Parameter | Value |
| --- | --- |
| Gear ratio | 24.65:1 |
| Motor full steps/revolution | 200 |
| Microsteps/full step | 256 in the ultra-slow configuration |
| Pulses/table revolution | 1,262,080 |
| Pulses/table degree | 3,505.78 |
| Command increment | ~0.000285°/pulse |
| Encoder CPR | 5,000 |
| Effective table counts/revolution | 246,500 |
| Measured increment | ~0.001460°/count |

# 9. Physical Motion Evidence

Figure 12. Representative rotation-rate demonstration under multiple load and offset conditions.

The physical platform demonstrated commanded rotation, lateral offset, encoder feedback, and load support. Configuration-specific speed limits are documented in the companion firmware report rather than treated as one universal mechanical range.

EVIDENCE SCOPE The physical evidence supports prototype operation and integration. It does not constitute formal production, EMC, environmental, or downstream optical-system qualification.

# 10. Requirements and Verification Status

| ID | Requirement | Status / evidence |
| --- | --- | --- |
| MECH-01 | Support the optical assembly and applied static load | 40 lb demonstration |
| MECH-02 | Provide lateral adjustment | More than ±3 in travel |
| MECH-03 | Provide operator-readable translation | Fixed scale and high-contrast pointer |
| ELEC-01 | Receive differential encoder A/B | AM26LS32 interface schematic and PCB |
| ELEC-02 | Route STEP/DIR control | Controller and R725 interface documented |
| ELEC-03 | Provide deliberate ground reference | Solid-copper plane and common ground |
| SAFE-01 | Provide independent stop behavior | Emergency stop and target deadman |

# 11. Limitations and Release Actions

| Item | Controlled statement |
| --- | --- |
| Prototype manufacture | No formal shock, vibration, thermal, EMC, or long-duration certification. |
| Electrical power | Idle, steady-state, and peak system power remain to be measured. |
| Fabrication release | Editable board files and final manufacturing package remain controlled. |
| Integrated build | Preserve one tagged firmware/hardware release for configuration traceability. |
| Metrology | Motion-platform characterization does not certify downstream optical performance. |

# 12. Attribution and Revision Control

Herder Elektronische Systemen developed the Dual-Core Optical Motion Testbed platform architecture, custom encoder/STEP-DIR PCB, dual-core firmware integration, host software, test workflows, and customer-facing documentation represented here. No third-party affiliation or endorsement is claimed.

| Rev | Date | Description |
| --- | --- | --- |
| 2.9 | 2026-08-04 | Renamed the public project and rewrote the platform report identity while retaining the exploded assembly, schematic, routed copper, solid-copper ground plane, PCB renders, and physical verification evidence. |